## Slide 1: Why Use Mass Spectrometry to Measure Proteins?
Problems with clinical immunoassays
Single-plex
Poor standardization
Hook effect
Autoantibodies
Hoofnagle and Wener, J Immunol Methods (2009)

## Slide 2
Traditional Automated Immunoassays
Capture antibody selects sandwiches detected by enzyme-conjugated secondary Generally, one at a time

## Slide 3
Y
Y
Y
Y
Y
Y
Can We Multiplex Immunoassays?
vs.

## Slide 4
Not easily.
Can We Multiplex Immunoassays?

## Slide 5
Hook Effect
Too much tumor antigen can lead to negative result

## Slide 6
Autoantibody Interference
Autoantibodies can sequester antigen away from detectable complexes
Falsely negative results with sandwich assays

## Slide 7: Mass Spectrometry: A Potential Solution
Direct detection of analyte by mass spectrometry
Interlaboratory calibration
Improved specificity
No sandwiches formed
No hook effect
Saturation of signal – plateau of calibration curve
Digestion of all proteins
Destroy autoantibodies

## Slide 8: Discovery Proteomics: Is it suitable to replace immunoassays?

## Slide 9: Shotgun Proteomics
µLC
Protein Mixture
proteolysis
Peptide Mixture
MS
MS/MS
1
1
2
2
3
3

## Slide 10: These Mixtures are Complicated!
Time (Min)
m/z
MS Scan
MS/MS
MS/MS
MS/MS

## Slide 11: These Mixtures are Complicated!
Time (Min)
m/z

## Slide 12: Sampling Every Precursor Using DDA
To get every precursor using DDA, we would have to collect 121 MS/MS spectra on average per chromatographic peak
200
150
50
100
0
Scan index
300
250
Number of precursors from m/z 400 to 1,000
Number of precursors
Average:
120.58
Notes: We count the number of precursor in each scan

## Slide 13: Data Dependent Acquisition will Always Under Sample

## Slide 14: DDA always misses and is least reproducible on the low abundant signals
Hoopmann et al, JPR 2009

## Slide 15: Visiting Seattle from Tallest to Smallest
2) Washington Mutual Tower
3) Two Union Square
1) Columbia Center
4) Seattle Municipal Tower
5) F5 Tower
6) Safeco Plaza
7) US Bank Center
8) Space Needle
9) Russell Investments Center
10) DocuSign Tower
11) 800 Fifth Ave
12) 901 Fifth Ave

## Slide 16: You would miss …
If you wouldn’t visit Seattle by going to the tallest structures, why would we use this approach to understand protein mixtures???

## Slide 17: Mass Spectrometry Data Acquisition Strategies Used in Proteomics
Data Dependent Acquisition (DDA)
Peptide Identification
NLPFSVENK
Database Searching

## Slide 18: So what do we do?
Build on Other People’s Knowledge

## Slide 19: So what do we do?

## Slide 20: So what do we do?

## Slide 21: C. elegans as an In Vivo Model System for Proteomics
The only multicellular organism with a genome sequenced to completeness
Complete Lineage for all 959 Somatic Cells

## Slide 22
Insulin Receptor
DAF-2
IRS1-4
IST-1
pi 3-kinase
AGE-1
PDK1
PDK-1
AKT/PKB
AKT-1/AKT-2
FKHRL1
DAF-16
FKHR
AFX
Humans
C. elegans

## Slide 23: Fractionation to Improve the Coverage of Proteins Involved in Insulin / IGF-1 Signaling
Soluble Proteins
Insoluble Proteins
Density Centrifugation 10 Fractions
Reversed Phase 10 Fractions
(NH4)2SO4 Cuts 8 Fractions
10 MudPITs
8 MudPITs
10 MudPITs
>6,000,000 MS/MS Spectra, 67,047 Unique Peptide Identifications, and 6,779 Unique Protein Identifications
0 Proteins Identified that are Known to be Involved in the Insulin / IGF-1 Signaling Pathway

## Slide 24: Mass Spectrometry Data Acquisition Strategies Used in Proteomics
Data Dependent Acquisition (DDA)
Peptide Identification
NLPFSVENK
Database Searching
Parallel Reaction Monitoring (PRM)
Peptide Quantification
NLPFSVENK
Targeted Analysis
Cycle time

## Slide 25
PRM
SRM
Targeted Acquisition
Targeted Extraction
Mass Analysis
Mass Filter
Unit Resolution
Full Spectrum
Often Accurate mass

## Slide 26
Ion Source
CID
521.7
757.6
Q1
Q3

## Slide 27
Precursor Ion > Product Ion
TGASEAVPSEGK > GK
TGSAEAVPSEGK > EGK
TGSAEAVPSEGK > SEGK
TGSAEAVPSEGK > PSEGK
TGSAEAVPSEGK > VPSEGK
TGSAEAVPSEGK > AVPSEGK
TGSAEAVPSEGK > EAVPSEGK
TGSAEAVPSEGK > AEAVPSEGK
TGSAEAVPSEGK > SAEAVPSEGK
566.78 > 204.13
566.78 > 333.18
566.78 > 420.21
566.78 > 517.26
566.78 > 616.33
566.78 > 687.37
566.78 > 816.41
566.78 > 903.44
566.78 > 1031.5
y2
y3
y4
y5
y6
y7
y8
y9
y10
SRM Chromatograms

## Slide 28
Synthetic Peptide Standards
Worm Homogenate 25 µg
75 attomole IDATTHIGGVQIK
7.5 femtomole QEDVVIEGWLHK
IDATTHIGGVQIK DAF-16
QEDVVIEGWLHK AKT-1

## Slide 29: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 30: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 31: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 32: 12 Attomoles of ILFDYSK on Column in 10 µg of Unfractionated Human Plasma
*

## Slide 33: Questions When Developing a Targeted Assay
What peptides are 1) detectable, 2) chromatograph well, 3) are unique to, and 4) are stable and robust over time.
What transitions provide a sensitive and unique measure of that peptide

## Slide 34
A Case Study in Developing a Targeted “Wellness Assay” from Dried Blood Spots by Empirical Refinement

## Slide 35: Selection of Protein Targets for DBS Wellness Assay
Adapted from: Tables 1& 2 - N. Leigh Anderson. Clin. Chem. (2010), 56, 177-185.
Acid Phosphatase
Alanine Aminotransferase
Albumin
Aldolase
Alkaline Phosphatase
α-1-acid glycoprotein
α -1-antitrypsin (SERPINA1)
α -1-antichymotrypsin (SERPINA3)*
α -1-antiplasmin (SERPINF2)
α -2-HS-glycoprotein (FetuinA)
α –fetoprotein (AFP)
Amylase
ACE
Antithrombin III (SERPINC1)
Apolipoprotein A1
Apolipoprotein B
Apolipoprotein C2*
Apolipoprotein C3*
Apolipoprotein E*
Apolipoprotein H*
Apolipoprotein J*
Aspartate Aminotransferase
β -2-microglobulin
β -Thromboglobulin
Biotinidase
Cancer antigen 125
Cancer antigen 15-3
Human epididymis protein 4
Carcinoembryonic antigen (CEA)
Ceruloplasmin
Cholinesterase
Complement C1 (C1R/C1S)
Complement C1 inhib.
Complement C1Q
Complement C3
Complement C4
Complement C5
C-reactive Protein
Creatine Kinase-BB
Creatine Kinase-MM
Cystatin C
Erythopoeitin
Factor IX antigen
Factor X
Factor XIII
Ferritin
Fibrinogen
Fibronectin
FSH
 glutamyl transferase (I/V)
Haptoglobin
β-hCG
Hemoglobin A1C (HBB)
Hemopexin
Her-2/neu
Human growth hormone
Human placental lactogen
IgA
IgD
IgE
IgG
IgM
Inhibin-A
Insulin
IGF-1
IGF-2
IGFBP-1
IGFBP-2
Interleukin-2 receptor
Isocitric dehydrogenase
light chains
Kininogen 1*
Lactate Dehydrogenase (heart)
Lactoferrin
 light chains
Lipase
Lp(a)
LP-PLA2
Leutinizing hormone
Lysozyme
Myeloperoxidase
Myoglobin
Osteocalcin
Parathyroid hormone
Phosphohexose isomerase (GPI)
Plasminogen
Plasminogen activator inhib.
Prealbumin (transthyretin)
BNP
Procalcitonin
Prolactin
Properidin Factor B (CFB)
Prostatic acid phosphatase
PSA
Protein C
Protein S
Pseudocholinesterase
Pyruvate kinase
Renin
Retinol binding protein
Mesothelin-related peptide
Sorbital dehydrogenase
Thyroglobulin
TSH
Thyroxine binding globulin
Tissue plasminogen activator
Transferrin
Troponin T (TnT)
Troponin I (TnI)
Trypsin
Urokinase
Vitamin D binding protein*
Vitronectin
Von Willebrand factor
Zinc- α -2-glycoprotein*

## Slide 36: Spectrum Counts vs. SRM Signal Intensity: Apolipoprotein B100
Chart: SRM Signal Intensity
# of +2 charge state spectra in 12 independent DDA runs | 7 | 5 | 2 | 4 | 5 | 4 | 7 | 3 | 2 | 8 | 2 | 7 | 5 | 5 | 1 | 2 | 7 | 7 | 4 | 4 | 6 | 3 | 2 | 2 | 6 | 3 | 1 | 3 | 2 | 1 | 10
Prakash, A. et al. J Proteome Res. 2009 8(6): 2733–2739.
Notes: Although these approaches are better than selecting proteotypic peptides at random, our results suggest that current predictor algorithms and spectral counting approaches provide imperfect ranking and identification of optimal proteotypic peptides, potentially limiting the utility of large-scale peptide-synthesis efforts that rely on such approaches as a first-round filter.

## Slide 37: Our Refinement Workflow
Method Development
Process
Skyline
Matrix
Peptide: 7-23 aa, +2
Product: from y3 to yn − 1, +1
Protein Extraction
Digestion
Triple quadrupole
SRM analysis
Data analysis
Recombinant protein production
Monitor all tryptic peptides
Selection of transitions
and iRT
Peptide stability analysis
Validation of transitions in the matrix
Optimization of transitions
Filter and selection
4 transitions/peptide
m/z fragment > m/z precursor
Filter and Selection
Build library in Panorama
Recombinant Library (Dotp) and Retention time prediction
Filter and Selection
3 peptides per protein
RT standards;
GST heavy
CFB heavy
Control
n=8
OSCC
n=22
Number of peptides/ protein
3
~30
~20
~10
~5
10 µg

## Slide 38
Workup of Recombinant/Native Protein Standard
http://dnasu.org
Gene Product of Interest
GST
IEAIPQIDK
IEAIPQIDK
+
Expand, Purify, & Sequence cDNA clone
In Vitro Transcription and Translation
Enrichment via Glutathione-Sepharose Resin
~25-500 femtomoles/25 uL IVT
< $20.00/Protein
Notes: The 1-Step Human IVT System uses HeLa cell lysates to perform cell-free protein expression. Using a human-based system yields proteins that mirror natural proteins with respect to post-translational modifications and results in proteins that exhibit natural protein function and structure.

## Slide 39: Our Refinement Workflow
Method Development
Process
Skyline
Matrix
Peptide: 7-23 aa, +2
Product: from y3 to yn − 1, +1
Protein Extraction
Digestion
Triple quadrupole
SRM analysis
Data analysis
Recombinant protein production
Monitor all tryptic peptides
Selection of transitions
and iRT
Peptide stability analysis
Validation of transitions in the matrix
Optimization of transitions
Filter and selection
4 transitions/peptide
m/z fragment > m/z precursor
Filter and Selection
Build library in Panorama
Recombinant Library (Dotp) and Retention time prediction
Filter and Selection
3 peptides per protein
RT standards;
GST heavy
CFB heavy
Control
n=8
OSCC
n=22
Number of peptides/ protein
3
~30
~20
~10
~5
10 µg

## Slide 40: nanoLC MS/MS: SRM Signal Intensity Rank
Notes: Feel for peak shape

## Slide 41: Our Refinement Workflow
Method Development
Process
Skyline
Matrix
Peptide: 7-23 aa, +2
Product: from y3 to yn − 1, +1
Protein Extraction
Digestion
Triple quadrupole
SRM analysis
Data analysis
Recombinant protein production
Monitor all tryptic peptides
Selection of transitions
and iRT
Peptide stability analysis
Validation of transitions in the matrix
Optimization of transitions
Filter and selection
4 transitions/peptide
m/z fragment > m/z precursor
Filter and Selection
Build library in Panorama
Recombinant Library (Dotp) and Retention time prediction
Filter and Selection
3 peptides per protein
RT standards;
GST heavy
CFB heavy
Control
n=8
OSCC
n=22
Number of peptides/ protein
3
~30
~20
~10
~5
10 µg

## Slide 42: nanoLC MS/MS: Fragmentation Fingerprint
R.TGISPLALIK.G [219, 228]
K.ALVEQGFTVPEIK.T [2578, 2590]
K.TTLTAFGFASADLIEIGLEGK.G [674, 694]
R.YEDGTLSLTSTSDLQSGIIK.N [1536, 1555]

## Slide 43: iRT Calibration

## Slide 44: Our Refinement Workflow
Method Development
Process
Skyline
Matrix
Peptide: 7-23 aa, +2
Product: from y3 to yn − 1, +1
Protein Extraction
Digestion
Triple quadrupole
SRM analysis
Data analysis
Recombinant protein production
Monitor all tryptic peptides
Selection of transitions
Peptide stability analysis
Validation of transitions in the matrix
Optimization of transitions
Filter and selection
4 transitions/peptide
m/z fragment > m/z precursor
Filter and Selection
Build library in Panorama
Recombinant Library (Dotp) and Retention time prediction
Filter and Selection
3 peptides per protein
RT standards;
GST heavy
CFB heavy
Control
n=8
OSCC
n=22
Number of peptides/ protein
3
~30
~20
~10
~5
10 µg

## Slide 45: nanoLC MS/MS: Peptide Stability Analysis
K.DNVFDGLVR.V [4168, 4176]
Standard
Standard
72hrs in A/S
R.TGISPLALIK.G [219, 228]
Standard
Standard
72hrs in A/S
K.TTLTAFGFASADLIEIGLEGK.G [674, 694]
Standard
Standard
72hrs in A/S

## Slide 46: Our Refinement Workflow
Method Development
Process
Skyline
Matrix
Peptide: 7-23 aa, +2
Product: from y3 to yn − 1, +1
Protein Extraction
Digestion
Triple quadrupole
SRM analysis
Data analysis
Recombinant protein production
Monitor all tryptic peptides
Selection of transitions
Peptide stability analysis
Validation of transitions in the matrix
Optimization of transitions
Filter and selection
4 transitions/peptide
m/z fragment > m/z precursor
Filter and Selection
Build library in Panorama
Recombinant Library (Dotp) and Retention time prediction
Filter and Selection
3 peptides per protein
RT standards;
GST heavy
CFB heavy
Control
n=8
OSCC
n=22
Number of peptides/ protein
3
~30
~20
~10
~5
10 µg

## Slide 47
Validation of Peptides/Transitions in Matrix
K.SPAFTDLHLR.Y [3979, 3988]
R.EFQVPTFTIPK.L [1333, 1343]

## Slide 48: Digestion Time Course
K.ALVEQGFTVPEIK.T [2578, 2590]

## Slide 49: 255 Peptides Measured in a Single LC-MS/MS Run (63 Proteins)

## Slide 50: 255 Peptides Measured in a Single LC-MS/MS Run (63 Proteins)

## Slide 51: Quantitative Analysis 101
What is does it mean to be quantitative?
Notes: Poll 3 … what does it mean to be quantitative

## Slide 52: Peak area measurements should scale with the amount of peptide in the sample
known quantity

## Slide 53: Quantitative Analysis 101
Amount (moles)
Measured Intensity (Area)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 54: Quantitative Analysis 101: Using an Internal Std
Amount (moles)
Intensity Ratio (Area Ratio)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 55: Quantitative Analysis 101
What is does it mean to be quantitative?
What does signal calibration mean?
Notes: Poll 4
What does signal calibration mean

## Slide 56: Measurements on different platforms are not measured on the same scale

## Slide 57: Between Batch/Platform/Lab Signal Calibration
Purpose of signal calibration
Placing signal on the same scale
Often the scale is arbitrary
  Kilogram = mass of a liter of water at freezing point
  Meter = 1/10,000,000 the distance from the equator to the north pole
  Pee dee belomnite (PDB) = reference for 13C, fossil found in Peedee Formation in South Carolina
  Standard mean ocean water (SMOW) = reference for 18O, arbitrary mix of distilled water from different regions of the world.

## Slide 58: Questions:
Can a method be quantitative if you don’t know the range over which the signal reflects the quantity?

## Slide 59: What does it mean if these are the two points?
Amount (moles)
Measured Intensity (Area)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 60: Questions:
Can a method be quantitative if you don’t know the range over which the signal reflects the quantity?
Design an approach to confirm linearity and LLOQ for peptides in a complex matrix?
  Assume you have < 50 peptides in your assay.
  Assume you have >1000 peptides in your assay?

## Slide 61: Matched Matrix Dilution Curves
Building on: Grant and Hoofnagle, Clin Chem 2014
Pino et al, JPR 2020

## Slide 62: Not all Peptides are Equal Quantitative Proxies for a Protein
Pino et al, JPR 2020
Notes: plasma membrane ATPase (PMA1

## Slide 63: Two other ways we can think of
Take two samples with extremes in the signal measured from a discovery experiment
  Mix samples in different ratios and demonstrate that there is a linear response between the extremes
Method of standard addition

## Slide 64: A key reminder: We measure peptides and not proteins.
Plubell et al, BioRxiv 2021

## Slide 65: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 66: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 67: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 68
Protein-level quantification
Example: GAPDH (glyceraldehyde-3-phosphate dehydrogenase) in cerebrospinal fluid
Peptide level quant can reflect protein level differences
Notes: So here we have a protein from this cerebrospinal fluid, GAPDH, and for the protein quantification we are assuming that all the peptides that map to the protein coding sequence and used for generating that value will be increased in the Parkinson’s compared to the other two groups,

## Slide 69
Peptide-level quantification
Protein-level quantification
Example: GAPDH (glyceraldehyde-3-phosphate dehydrogenase) in cerebrospinal fluid
Peptide level quant can reflect protein level differences
Notes: In this case all our peptides have the same trend across the experimental groups, and based on this data it is probably fine to aggregate these peptides to a single value.

## Slide 70
Example: SCG2 (secretogranin-2) in cerebrospinal fluid
Peptide-level quantification
Protein-level quantification
Summing to protein-level can mask interesting differences
Notes: However, we have cases in which the protein quant doesn’t seem to be changing across experimental groups, but on the individual peptide level we have some peptides that do appear to change, sometimes in opposite directions from one another.

## Slide 71
Ways to improve the quantitative precision

## Slide 72: Metabolically Labeling as an Internal Standard
15N Enriched Media
Protein Synthesis
15N Internal Standard
Sample
Cell Lysis
Reduction and Alkylation
Digestion
µLC/µLC/MS/MS
Sample & Int Std Combined

## Slide 73
Identify Peptide
Measure Ratio A
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Proteolytic Digestion
Cell Lysis
µLC-MS/MS
Identify Peptide
Measure Ratio B
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Derivatization
Cell Lysis
µLC-MS/MS
Proteolytic Digestion

## Slide 74
Identify Peptide
Measure Ratio A
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Proteolytic Digestion
Cell Lysis
µLC-MS/MS
Identify Peptide
Measure Ratio B
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Derivatization
Cell Lysis
µLC-MS/MS
Proteolytic Digestion

## Slide 75: There are lots of different ways to incorporate an internal standard