## Slide 1: What is signal?

## Slide 2: What is signal?
Time-varying voltage, current or electromagnetic waves that carry information
Analog or digital representations of analog physical quantities
Analog signal: a continuous value in which a time varying quantity (i.e. voltage, pressure, etc…) represents another time-based variable – one variable is an “analog” of the other.
  Dimmer switch connected to a light bulb. The dimmer has an infinite number of positions from off and full – and a corresponding infinite number of output from the light bulb
Digital signal: express the variation in discrete values.
  More like an “on/off” switch.
Most analog signal is converted to a digital form so that it can be transmitted, stored, or processed.

## Slide 3: Simple Transducer Circuit
Signal
1 2 3 4 5 6
Tranducer
Amplifier
Analog to Digital Converter
Readout
Current = i
k*i
v = f(i)
Number

## Slide 4: Transducers are similar for photons and ions
Photomultiplier Tube
Electron Multiplier

## Slide 5: Important aspects of ADCs (digital conversion)
Sampling Rate
  How fast the ADC can convert the signal?
Bit resolution
  How much accuracy can the ADC convert from analog to digital?
3-bit ADC can provide 23 divisions (8) in signal
8-bit ADC = 256 division
16-bit ADC = 65536

## Slide 6: ADC Bit Resolution
https://claude.site/artifacts/47f9c1a3-6103-4c80-8746-e5797836833c

## Slide 7: ADC Bit Resolution
https://claude.site/artifacts/47f9c1a3-6103-4c80-8746-e5797836833c

## Slide 8: ADC Bit Resolution

## Slide 9: What is noise?

## Slide 10: Types of Noise
Johnson noise (aka Thermal noise): noise from the thermal motion of electrons in the electronics. Constant and independent of the measured signal. Dependent on temperature and resistance in the electronics. Can be limiting factor at very high sensitivity measurements. Homoscedastic
Shot noise (aka Poisson noise, counting noise): Random variations in electron or ion motion lead to random variations in current. The higher the current, the larger the shot noise. Heteroscedastic

## Slide 11: Types of Noise
Heteroscedastic noise – Signal dependent noise
https://claude.site/artifacts/224c911e-6b54-42bf-9804-114fa3104d3e

## Slide 12: Types of Noise
Heteroscedastic noise – Signal dependent noise
https://claude.site/artifacts/224c911e-6b54-42bf-9804-114fa3104d3e

## Slide 13: Types of Noise
Homoscedastic noise – Signal independent noise
https://claude.site/artifacts/224c911e-6b54-42bf-9804-114fa3104d3e

## Slide 14: Types of Noise
Homoscedastic noise – Signal independent noise
https://claude.site/artifacts/224c911e-6b54-42bf-9804-114fa3104d3e

## Slide 15: Both exist at the same time but from unknown amounts

## Slide 16: Shot Noise
~10x Improvement in Signal
~3x Improvement in S/N

## Slide 17: Ways to express Signal to Noise?
S = height of the peak above baseline
= baseline
Noise = Standard deviation of the baseline
Most Common Way

## Slide 18: Common way to calculate S/N of a peak

## Slide 19: Ways to express Signal to Noise?
S = height of the peak above baseline
= baseline
Noise = Stdev of the baseline
Most Common Way
What is wrong with this?
Signal to Baseline Noise

## Slide 20: Propagation of Error
Type | Example | Standard Deviation
Addition or Subtraction |  | 
Multiplication or Division |  | 
Exponential |  | 
Logarithmic |  | 
Anti-Logarithmic |  | 
Type | Example | 
Addition or Subtraction |  | 
Multiplication or Division |  | 
Exponential |  | 
Logarithmic |  | 
Anti-Logarithmic |  | 
Where a, b, c are measured variables from an experiment and , , and are the standard deviations.

## Slide 21: What sort of noise source is dominant?

## Slide 22: Glu1-fib peptide fragmentation pattern
y6
y3
y4
y7
y8
y9
y10
y11

## Slide 23: Calculating the ratio of highest y-ion intensities to y9 ion
y6
y3
y4
y7
y8
y9
y10
y11
Example: y7-to-y9
1.247

## Slide 24: Calculate ratios across 2,000 replicate spectra
Using an intra-spectrum ratio minimizes non-poisson noise effects (i.e. electrospray)

## Slide 25: Derivation
Peak area of fragment ion
Fit parameter
Number of ions
Injection time
Ratio between two signals
and t are constants
Propagation of errors to calculate the error in the ratio
MacCoss et al, Anal Chem (2001)
Canterbury, JASMS (2014)
Schwartz, Patent 7,109,474 (2004)

## Slide 26: Derivation
total variance in a repeated measurement
variance due to Poisson (shot) noise
variance due to all other noise sources
Simplify, assuming
Peak area equation
MacCoss et al, Anal Chem (2001)
Canterbury, JASMS (2014)
Schwartz, Patent 7,109,474 (2004)

## Slide 27: Formula to determine calibration (ions/sec)
variance in the ratio of peak areas
ratio of peak areas
A = area of fragment peak
t = injection time
Calibration factor
Signal Independent noise
MacCoss et al, Anal Chem (2001)
Canterbury, JASMS (2014)

## Slide 28: Calibration plots for Astral, Stellar, and Orbitrap
Chris Hsu

## Slide 29: Other sources of noise?
Chemical noise (aka interference)
  Interactions with other molecules in:
    Sample
    Room
Environmental noise
  Frequencies that are picked up by RF detection electronics

## Slide 30: Common Examples:
Dolby noise reduction: perform FFT, remove high frequency portion of the spectrum, inverse FFT
Noise cancelling headphones: A microphone records what is happening outside and the exact opposite frequency is applied to cancel those sounds out.

## Slide 31: Precision vs Accuracy
Precision is about taking multiple attempts—such as shots, guesses, or measurements—and having all of them end up very close to each other.
Accuracy is about having your attempts land at the true or desired location.

## Slide 32: Problem:
You have two methods for making a measurement and they give different results.
How do you assess if there is a bias?
  Any systematic error that could distort the accuracy of the research findings
Sometimes a new method is developed, and we need to compare the new method with the existing one regarded as a “reference method”.
Frequently, however, we cannot regard either method as giving the true value of the quantity being measured
"Do the two methods of measurement agree sufficiently closely?"

## Slide 33: What does this tell you?
Altman and Bland, The Statistician 1983

## Slide 34: Bland-Altman Plot
Altman and Bland, The Statistician 1983

## Slide 35: Assess if there is a bias
Left ventricular ejection fraction plotted to Illustrate Bias

## Slide 36: Measurement Error Simulation
https://claude.site/artifacts/2e700dd0-43c8-4b19-9be7-971e9d190360
Notes: With Bias but equal random error

## Slide 37: Measurement Error Simulation
Notes: No systematic error. Differences in the Random error (this last part isn’t visible.