## Slide 1: Mass Spectrometry 101 Ionization and Analyzers

## Slide 2: Basic Features of a Mass Spectrometer
PC
Work Station
Ion
Source
Mass
Analyzer
Detector
“High” Vacuum
Answers
Out
Sample
In
You

## Slide 3
“I feel sure that there are many problems in chemistry which could by solved with far greater ease by this than any other method. The method is surprisingly sensitive - more so than that of spectrum analysis - requires an infinitesimal amount of material and does not require this to be specially purified...”
Thomson, 1913
Biology/Medicine

## Slide 4: Nobel Prizes in Mass Spectrometry
J.J. Thomson (1906 in Physics) – Cambridge, UK
  For the theoretical and experimental investigations of the behavior of gas phase ions in a magnetic fields (e.g. the mass spectrometer)
H. Dehmelt
W. Paul
J.J. Thomson
Francis William Aston (1922 in Chemistry) – Cambridge, UK
  For his discovery, by means of his mass spectrograph, of isotopes
Hans Dehmelt and Wolfgang Paul (1989 in Physics) – UW, USA and University of Bonn, Germany
  For the development of the quadrupole mass filter and the quadrupole ion trap
F.W. Aston

## Slide 5: Nobel Prizes in Mass Spectrometry
John Fenn – Electrospray Ionization
2002 Nobel Prize in Chemistry
“for the development of methods for identification and structure analyses of biological macromolecules”
“for their development of soft desorption ionization methods for mass spectrometric analyses of biological macromolecules “
Koichi Tanaka – Laser Desorption Ionization

## Slide 6: Mechanism of Electrospray Ionization (ESI)
+
-
+
+
+
+
+
+
+
+
+
+
-
-
-
+
+
+
+
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Charged
Capillary
Needle
Heated
Capillary
+3-8 kV
Taylor cone
Rayleigh
limit hit:
Coulombic
explosion
+
+
+
+
+
+
+

## Slide 9: Nanospray: How Sensitivity is Obtained
Improved Droplet Collection Efficiency
Smaller Droplets = More Surface Area
d ≈ k.(flow)2/3 (Mann-Wilm Theory)
Tang, Smith Acc. Chem. Res. 2004, 37, 269

## Slide 10: How Does Nanospray Yield Sensitivity?
From Tang & Smith, Acc. Chem. Res. 2004, 37, 269-278

## Slide 11
160 nL/min
1000 nL/min
300 nL/min
500 nL/min
10 µm Emitter, 50% ACN (0.1% FA)
Nanospray:Emission from a Liquid “Tip”

## Slide 12: ESI-MS Spectrum of Horse Apomyoglobin
Apomyoglobin
MW=16,951.2
+18
+20
+16
+14
+12
+10
+22
m/z

## Slide 13: ESI-MS Spectrum of Bovine Serum Albumin
Albumin
MW=66,621
+41
+52
+34
m/z

## Slide 14
m/z
Where:
zi = charge of peak Pi
Pj = peak with charge zi + 1

## Slide 15: Quadrupole Mass Filter
-
-
+
+
U + V cos(t)

## Slide 16: Quadrupole Mass Filter
The high pass rods filter out ions
with too low of a m/z
The low pass rods filter
out ions with too high of
a m/z
A combination of DC and
RF potential is applied
and varied with time to
separate ions
low pass (-)
high pass (+)

## Slide 17: How a quadrupole mass filter works
+ DC Voltage will focus + ions on Z axis
The negative portion of an RF V will have a larger effect on a low mass ion.
Result: Only high mass ions pass
- DC Voltage will draw + ions away from Z axis
The positive portion of an RF V will have a larger effect on a low mass ion steering it along the Z axis.
Result: Only low mass ions pass
Combination of the high pass and low pass filters creates a “notch” where a narrow mass can pass.
Notes: If an ion is very heavy and/or the frequency of the RF potential is very rapid, an ion will tend only to feel the effect of the average potential applied to the electrode structure. That is, heavy ions will tend only to be influenced by the positive dc potential. The small time when the electrodes are negative will have negligible effect on heavy ions because of the + DC field
With a negative DC voltage, a positive ion is being drawn to the electrodes. A light ion will respond better to the positive portions of the RF field.

## Slide 18: Tandem Mass Spectrometry (MS/MS)
ESI
MS1
MS2
CID
from CID+MS2
from MS1
400
600
800
1000
1200
1400
1600
1800
m/z
0
20
40
60
80
100
Relative Abundance
Tandem In Space

## Slide 19: Tandem Mass Spectrometry (MS/MS) Analysis Peptides
Ion Source
MS1
MS2
CID
Data System
y1
y2
b1
b2

## Slide 20
S Q H L F A P I S E K
Notes: How do we derive a peptide sequence from it’s MS/MS spectrum. Predominantly amino acids are fragmented at amide bonds and because amino acids have different side chains we can the amino acid sequence from the difference in mass in adjacent fragments.

## Slide 21
S Q H L F A P I S E K
147
276
363
477
574
645
792
905
1042
1170
216
353
467
614
685
782
895
982
1111
88
1111
1042
905
982
792
895
685
644
M-NH3
574
1042 / 2
467
353
Notes: When you look at all the spectra data together you can deduce the sequence of this peptide ion. You can see in red the masses of the peptides fragmented between the Leu and Phe residues.

## Slide 23: Modern ion traps make use of mass selective instability
Stafford, JASMS 2002

## Slide 24: Acquiring Mass Spectrometry Data of Peptides
Peptides from the column are ionized and transmitted into the vacuum system using electrospray and trapped in the 2-dimensional ion trap
Microcapillary Column
2 kV
30 V

## Slide 25: Acquiring Mass Spectrometry Data of Peptides
m/z
Abundance
Peptides from the column are ionized and transmitted into the vacuum system using electrospray and trapped in the 2-dimensional ion trap
Ions are selectively ejected radially by their m/z to produce a mass spectrum
The instrument data-system automatically selects a m/z for further analysis
Selected m/z

## Slide 26: Acquiring a Fragmentation Spectrum
The trap is then refilled

## Slide 27: Acquiring a Fragmentation Spectrum
The trap is then refilled
All ions are ejected from the trap except the predetermined m/z
The energy of the isolated ion is increased resulting in collisions with the helium gas in the trap

## Slide 28: Acquiring a Fragmentation Spectrum
m/z
Abundance
The collisions with the helium molecules result in structure specific fragment ions
The fragment ions are then selectively ejected to produce a tandem mass spectrum

## Slide 29: Matrix Assisted Laser Desorption/Ionization (MALDI)
Sample / Matrix
25 kV
N2 Laser – 337 nm

## Slide 30: Matrix Assisted Laser Desorption/Ionization (MALDI)
Sample / Matrix
25 kV
N2 Laser – 337 nm

## Slide 31: Matrix-Assisted Laser Desorption (MALDI) Time of Flight (TOF) Mass Spectrometry
Detector
337 nm N2 UV
Laser Pulse
25 kV
Vacuum

## Slide 33: Proteomics: Traditional Technology
Characterize peptides by mass spectrometry
N
C
C
N
2D Gel
1D Gel

## Slide 34: Peptide Mass Fingerprinting
MALDI
Detector
…R ASDEWK LHPSR WEQCPEFGR…
(2) Time of Flight
(1) Digest
protein
(3) Set up
mass list
609
735
1152
…
(4) Compare mass list against protein database of peptides