## Slide 1: Bottom-Up Proteomics
µLC
Protein Mixture
proteolysis
Peptide Mixture
MS
MS/MS
1
1
2
2
3
3

## Slide 2: Data Dependent Acquisition (DDA)
Mass Spectrum
Acquire mass spectrum
Select peptides for tandem mass spectrometry
Selectively isolate and fragment peptides.
Repeat for 3 different peptides
1
2
3
Tandem Mass Spectra of Peptides
~1 Second per cycle

## Slide 3
Automated Interpretation of MS/MS Spectra
MS/MS Spectrum
HTTLDEM
DTTEMIH
TMETIDH
Search Database for Precursor Mass
Theoretical MS/MS Spectra
Cross Correlation (Sequest)
Automated
Sequence and
Modification Identification
Eng et al. JASMS, 5, 976 (1994)

## Slide 4: Why Use Mass Spectrometry to Measure Proteins?
Problems with clinical immunoassays
Single-plex
Poor standardization
Hook effect
Autoantibodies
Hoofnagle and Wener, J Immunol Methods (2009)

## Slide 5
Traditional Automated Immunoassays
Capture antibody selects sandwiches detected by enzyme-conjugated secondary Generally, one at a time

## Slide 6
Y
Y
Y
Y
Y
Y
Can We Multiplex Immunoassays?
vs.

## Slide 7
Not easily.
Can We Multiplex Immunoassays?

## Slide 8
Hook Effect
Too much tumor antigen can lead to negative result

## Slide 9
Autoantibody Interference
Autoantibodies can sequester antigen away from detectable complexes
Falsely negative results with sandwich assays

## Slide 10: Mass Spectrometry: A Potential Solution
Direct detection of analyte by mass spectrometry
Interlaboratory calibration
Improved specificity
No sandwiches formed
No hook effect
Saturation of signal – plateau of calibration curve
Digestion of all proteins
Destroy autoantibodies

## Slide 11: Discovery Proteomics: Is it suitable to replace immunoassays?

## Slide 12: Bottom Up Proteomics
µLC
Protein Mixture
proteolysis
Peptide Mixture
MS
MS/MS
1
1
2
2
3
3

## Slide 13: These Mixtures are Complicated!
Time (Min)
m/z
MS Scan
MS/MS
MS/MS
MS/MS

## Slide 14: These Mixtures are Complicated!
Time (Min)
m/z

## Slide 15: Measure Signal Across Time
Schilling et al MCP 2012

## Slide 16: Measure signal versus time

## Slide 17: Peptide Quant 101
Amount (moles)
Measured Intensity (Area)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 18: Peptide Quant 101
Amount (moles)
Intensity Ratio (Area Ratio)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 19: Metabolically Labeling as an Internal Standard
15N Enriched Media
Protein Synthesis
15N Internal Standard
Sample
Cell Lysis
Reduction and Alkylation
Digestion
µLC/µLC/MS/MS
Sample & Int Std Combined

## Slide 20
Identify Peptide
Measure Ratio A
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Proteolytic Digestion
Cell Lysis
µLC-MS/MS
Identify Peptide
Measure Ratio B
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Derivatization
Cell Lysis
µLC-MS/MS
Proteolytic Digestion

## Slide 21
Identify Peptide
Measure Ratio A
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Proteolytic Digestion
Cell Lysis
µLC-MS/MS
Identify Peptide
Measure Ratio B
Normal Media
15N Enriched Media
Sample
Internal Std
Mix
Derivatization
Cell Lysis
µLC-MS/MS
Proteolytic Digestion

## Slide 22: There are lots of different ways to incorporate an internal standard

## Slide 23: Tandem Mass Tags (TMT)

## Slide 24: Sampling Every Precursor Using DDA
To get every precursor using DDA, we would have to collect 121 MS/MS spectra on average per chromatographic peak
200
150
50
100
0
Scan index
300
250
Number of precursors from m/z 400 to 1,000
Number of precursors
Average:
120.58
Notes: We count the number of precursor in each scan

## Slide 25: Data Dependent Acquisition will Always Under Sample

## Slide 26: DDA always misses and is least reproducible on the low abundant signals
Hoopmann et al, JPR 2009

## Slide 27: Visiting Seattle from Tallest to Smallest
2) Washington Mutual Tower
3) Two Union Square
1) Columbia Center
4) Seattle Municipal Tower
5) F5 Tower
6) Safeco Plaza
7) US Bank Center
8) Space Needle
9) Russell Investments Center
10) DocuSign Tower
11) 800 Fifth Ave
12) 901 Fifth Ave

## Slide 28: You would miss …
If you wouldn’t visit Seattle by going to the tallest structures, why would we use this approach to understand protein mixtures???

## Slide 29: Mass Spectrometry Data Acquisition Strategies Used in Proteomics
Data Dependent Acquisition (DDA)
Peptide Identification
NLPFSVENK
Database Searching

## Slide 30: So what do we do?
Build on Other People’s Knowledge

## Slide 31: So what do we do?

## Slide 32: So what do we do?

## Slide 33: C. elegans as an In Vivo Model System for Proteomics
The only multicellular organism with a genome sequenced to completeness
Complete Lineage for all 959 Somatic Cells

## Slide 34
Insulin Receptor
DAF-2
IRS1-4
IST-1
pi 3-kinase
AGE-1
PDK1
PDK-1
AKT/PKB
AKT-1/AKT-2
FKHRL1
DAF-16
FKHR
AFX
Humans
C. elegans

## Slide 35: Fractionation to Improve the Coverage of Proteins Involved in Insulin / IGF-1 Signaling
Soluble Proteins
Insoluble Proteins
Density Centrifugation 10 Fractions
Reversed Phase 10 Fractions
(NH4)2SO4 Cuts 8 Fractions
10 MudPITs
8 MudPITs
10 MudPITs
>6,000,000 MS/MS Spectra, 67,047 Unique Peptide Identifications, and 6,779 Unique Protein Identifications
0 Proteins Identified that are Known to be Involved in the Insulin / IGF-1 Signaling Pathway

## Slide 36: Mass Spectrometry Data Acquisition Strategies Used in Proteomics
Data Dependent Acquisition (DDA)
Peptide Identification
NLPFSVENK
Database Searching
Parallel Reaction Monitoring (PRM)
Peptide Quantification
NLPFSVENK
Targeted Analysis
Cycle time

## Slide 37
PRM
SRM
Targeted Acquisition
Targeted Extraction
Mass Analysis
Mass Filter
Unit Resolution
Full Spectrum
Often Accurate mass

## Slide 38
Ion Source
CID
521.7
757.6
Q1
Q3

## Slide 39
Precursor Ion > Product Ion
TGASEAVPSEGK > GK
TGSAEAVPSEGK > EGK
TGSAEAVPSEGK > SEGK
TGSAEAVPSEGK > PSEGK
TGSAEAVPSEGK > VPSEGK
TGSAEAVPSEGK > AVPSEGK
TGSAEAVPSEGK > EAVPSEGK
TGSAEAVPSEGK > AEAVPSEGK
TGSAEAVPSEGK > SAEAVPSEGK
566.78 > 204.13
566.78 > 333.18
566.78 > 420.21
566.78 > 517.26
566.78 > 616.33
566.78 > 687.37
566.78 > 816.41
566.78 > 903.44
566.78 > 1031.5
y2
y3
y4
y5
y6
y7
y8
y9
y10
SRM Chromatograms

## Slide 40
Synthetic Peptide Standards
Worm Homogenate 25 µg
75 attomole IDATTHIGGVQIK
7.5 femtomole QEDVVIEGWLHK
IDATTHIGGVQIK DAF-16
QEDVVIEGWLHK AKT-1

## Slide 41: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 42: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 43: Dilution of the Yeast Protein PGI1 in Undepleted Human Plasma
Peptide: ILFDYSK
R = 0.99992

## Slide 44: 12 Attomoles of ILFDYSK on Column in 10 µg of Unfractionated Human Plasma
*

## Slide 45: Questions When Developing a Targeted Assay
What peptides are 1) detectable, 2) chromatograph well, 3) are unique to, and 4) are stable and robust over time.
What transitions provide a sensitive and unique measure of that peptide

## Slide 46: Quantitative Analysis 101
What is does it mean to be quantitative?
Notes: Poll 3 … what does it mean to be quantitative

## Slide 47: Peak area measurements should scale with the amount of peptide in the sample
known quantity

## Slide 48: Assessment of LOD and LOQ
Definitions
Limit of Detection:
  Signal is statistically different from a blank
  Assumption is if a peptide is below the FDR in a peptide-centric search it is >LOD
Limit of Quantification:
  Point where a change in signal is reflective of the change in quantity
  Usually apply a precision threshold too
Differential Measurement:
  Two signals are statistically different
Pino et al, Anal Chem 2018

## Slide 49: Quantitative Analysis 101
What is does it mean to be quantitative?
What does signal calibration mean?
Notes: Poll 4
What does signal calibration mean

## Slide 50: Measurements on different platforms are not measured on the same scale

## Slide 51: Between Batch/Platform/Lab Signal Calibration
Purpose of signal calibration
Placing signal on the same scale
Often the scale is arbitrary
  Kilogram = mass of a liter of water at freezing point
  Meter = 1/10,000,000 the distance from the equator to the north pole
  Pee dee belomnite (PDB) = reference for 13C, fossil found in Peedee Formation in South Carolina
  Standard mean ocean water (SMOW) = reference for 18O, arbitrary mix of distilled water from different regions of the world.

## Slide 52: Questions:
Can a method be quantitative if you don’t know the range over which the signal reflects the quantity?

## Slide 53: What does it mean if these are the two points?
Amount (moles)
Measured Intensity (Area)
LOD
LLOQ
ULOQ
Linear Dynamic Range
LOD: Limit of Detection
LLOQ: Lower Limit of Quantitation
ULOQ: Upper Limit of Quantitation

## Slide 54: Questions:
Can a method be quantitative if you don’t know the range over which the signal reflects the quantity?
Design an approach to confirm linearity and LLOQ for peptides in a complex matrix?
  Assume you have < 50 peptides in your assay.
  Assume you have >1000 peptides in your assay?

## Slide 55: Matched Matrix Dilution Curves
Building on: Grant and Hoofnagle, Clin Chem 2014
Pino et al, JPR 2020

## Slide 56: Not all Peptides are Equal Quantitative Proxies for a Protein
Pino et al, JPR 2020
Notes: plasma membrane ATPase (PMA1

## Slide 57: Two other ways we can think of
Take two samples with extremes in the signal measured from a discovery experiment
  Mix samples in different ratios and demonstrate that there is a linear response between the extremes
Method of standard addition

## Slide 58: A key reminder: We measure peptides and not proteins.
Plubell et al, BioRxiv 2021

## Slide 59: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 60: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 61: Proteoforms complicate protein level measurements
Plubell et al, BioRxiv 2021
Notes: Take for example some theoretical protein as indicated on the left – by bottom up proteomics we can possibly detect 4 tryptic peptides that map to our database sequence.

## Slide 62
Protein-level quantification
Example: GAPDH (glyceraldehyde-3-phosphate dehydrogenase) in cerebrospinal fluid
Peptide level quant can reflect protein level differences
Notes: So here we have a protein from this cerebrospinal fluid, GAPDH, and for the protein quantification we are assuming that all the peptides that map to the protein coding sequence and used for generating that value will be increased in the Parkinson’s compared to the other two groups,

## Slide 63
Peptide-level quantification
Protein-level quantification
Example: GAPDH (glyceraldehyde-3-phosphate dehydrogenase) in cerebrospinal fluid
Peptide level quant can reflect protein level differences
Notes: In this case all our peptides have the same trend across the experimental groups, and based on this data it is probably fine to aggregate these peptides to a single value.

## Slide 64
Example: SCG2 (secretogranin-2) in cerebrospinal fluid
Peptide-level quantification
Protein-level quantification
Summing to protein-level can mask interesting differences
Notes: However, we have cases in which the protein quant doesn’t seem to be changing across experimental groups, but on the individual peptide level we have some peptides that do appear to change, sometimes in opposite directions from one another.