## Slide 1: Multispectral Imaging
Genome 575: Fundamentals of Biomedical Instrumentation
Brian Beliveau

## Slide 2
Recap from last time
Light in optical systems primarily comes from incandescence and fluorescence
Fluorescence relies on the excitation and decay of fundamental particles (usually e-)
The propagation of light can be altered by reflection, refraction, and absorption

## Slide 3
From Monday…
https://en.wikipedia.org/wiki/Jablonski_diagram#/media/File:Jablonski_Diagram_of_Fluorescence_Only-en.svg, https://www.expertsmind.com/topic/fluorimetry-and-phosphorimetry/jablonski-diagram-913753.aspx
Jablonski diagram
Heat
Kinetic
energy
DNA
Melanin

## Slide 4
From Monday, part 2
https://hubner-photonics.com/wp-content/uploads/2020/05/Performance-data-beam-profile-04-01-Series.png
Your camera chip (white) and Gaussian laser beam profile are shown ->
Is this set-up ideal for quantitative imaging? Why or why not?
Given the assumptions below, calculate the power density of the cyan box in kW/cm^2. How does this compare to the intensity of the sun?
  The central white + purple area is roughly uniform in intensity at 100 mW
  The cyan box is 1024 x 1024 100 nm pixels
Devise an engineering solution to make the illumination intensity uniform over the white box

## Slide 5
Solution 1: Refractive beam shaping
https://www.laserfocusworld.com/optics/article/14197298/flat-top-laser-beams-their-uses-and-benefits, https://pishaper.com/shaper.html

## Slide 6
Solution 2: Gradation ND filter
https://www.microscope.healthcare.nikon.com/products/photostimulation-tirf/ti2-lapp
Filter out
Filter in

## Slide 7
Picking your imaging bundle
https://media.newyorker.com/photos/59096d2b6552fa0be682fed7/master/pass/Wu-Dreaded-Bundle-Comes-to-Internet-TV1.jpg, https://m.media-amazon.com/images/I/51ZSG3TqyAS._AC_SL1500_.jpg
How many channels are in a typical microscope?
What bounds this number? (10 mins)

## Slide 8
Multiplexing challenges: light sources
https://en.wikipedia.org/wiki/File:Lasers.JPG , https://zeiss-campus.magnet.fsu.edu/articles/lightsources/images/metalhalidelampsfigure1.jpg, https://hubner-photonics.com/products/lasers/multi-line-lasers/
Most effective light sources bounded in UV-A, vis, near-IR
UV challenging to use with biological samples

## Slide 9
Multiplexing challenges: labels
https://www.idtdna.com/site/Catalog/Modifications/Dyes, http://www.chemspider.com/Chemical-Structure.23107060.html
Most effective dyes/labels concentrated in vis/near-IR
Rational design of organic dyes / FPs challenging
AlexaFluor 488

## Slide 10
Multiplexing challenges: optics
https://www.thorlabs.com/images/TabImages/MAP_A_Reflectance-780.gif, https://www.thorlabs.com/Images/GuideImages/1716_MatchedDoublet_6.jpg
Most optics can only transmit over defined ranges
𝝀-dependent differences lead to chromatic aberration
“Achromatic doublet”

## Slide 11
Multiplexing challenges: detectors
https://www.photometrics.com/wp-content/uploads/2019/09/QE-curve.png, https://www.photometrics.com/wp-content/uploads/2019/09/prime-bsi-shadow-750x696.png
CCD/sCMOS sensors most effective in vis
PMTs (laser scanning confocals/FACS) drop off a cliff in IR

## Slide 12
Multiplexing challenges: Stokes shifts
https://journals.plos.org/plosone/article/file?id=10.1371/journal.pone.0122848.g001&type=large, https://upload.wikimedia.org/wikipedia/commons/3/3f/Stokes_shift-_Rh6G.png
Stokes shift: max(Em) – max(Ex)
Width, spacing of Ex/Em peaks a huge factor in multiplexing

## Slide 13
Defining an imaging channel
https://encrypted-tbn0.gstatic.com/images?q=tbn:ANd9GcQGjmZByoM2TiwIz-P2benxkUGzlX3QJBgl2X7lUi5taLUpQLx0SmSb66f9A6v6PA0ui30&usqp=CAU, https://www.nikon.com/news/2016/img/pic_161109_01.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg
How is the microscope sub-setting the EM spectrum around a defined “channel”?
Light source (?), filter cube

## Slide 14
The anatomy of a filter cube
Ingredients: Ex filter (bandpass), DM (long-pass), Em fiter (bandpass)
Flip orientation 180º for inverted microscopes
https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg

## Slide 15
An example filter cube
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2

## Slide 16
An example filter cube: Ex filter
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
Bandpass
Transmit 450 – 490 nm

## Slide 17
An example filter cube: dichroic mirror
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
Longpass “beam splitter”
<495 nm: Reflect
≥495 nm: Transmit

## Slide 18
An example filter cube: Em filter
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
Bandpass
Transmit 500 – 550 nm

## Slide 19
An example filter cube
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
Excitation: 450 – 490 nm
Separation: 490 – 500 nm
Collection 500 – 550 nm

## Slide 20
Evaluating filter cubes
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
https://www.fpbase.org/spectra/

## Slide 21
Evaluating filter cubes
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
https://www.fpbase.org/spectra/

## Slide 22
Evaluating filter cubes
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/products/sets/49002-et-egfp-fitc-cy2
https://www.fpbase.org/spectra/
Em filter collection efficiency: 68.9%
Ex filter excitation efficiency: 75%

## Slide 23
Class activity 1: design a filter set (15 mins)
Assume you have the following laser lines: [405 nm, 488 nm, 532 nm, 561 nm, 594 nm, 640 nm]
Go to https://www.fpbase.org/spectra/
Select the dye “Cy3B” under the ‘Fluorophores’ tab and your choice(s) of laser in the ‘Light Sources’ tab
Design an “ideal” custom filter set (Ex, DM, Em) in the ‘Filters’ tab. You can substitute a long-pass “LP” filter for the DM (ie, only consider the transmission band)
Report your ‘Efficiency’ for Cy3B detection. You can set T% to 100 for all filters. Allow no excitation light to come through.

## Slide 24
Class activity 1: design a filter set (15 mins)
Em filter collection efficiency: 98.2%
Ex filter excitation efficiency: 59%

## Slide 25
Class activity 2: design a filter set pair (15 mins)
Assume you have the following laser lines: [405 nm, 488 nm, 532 nm, 561 nm, 594 nm, 640 nm]
Add the dye “Alexa Fluor 647” to the mix
Design an ideal filter set for Alexa 647 and if needed redesign your Cy3B filter set
Report your ‘Efficiency’ for Cy3B and Alexa Fluor 647 detection. You can set T% to 100 for all filters.

## Slide 26
Class activity 3: design a filter set trio (15 mins)
Assume you have the following laser lines: [405 nm, 488 nm, 532 nm, 561 nm, 594 nm, 640 nm]
Add the dye “Alexa Fluor 488” to the mix
Design an ideal filter set for Alexa 488 and if needed redesign your Cy3B and Alexa Fluor 647 filter sets
Report your ‘Efficiency’ for Cy3B, Alexa Fluor 647, and Alexa Fluor 488 detection. You can set T% to 100 for all filters.