## Slide 1: Objectives and PSFs
Genome 575: Fundamentals of Biomedical Instrumentation
Brian Beliveau

## Slide 2
Recap from last time
Commonly used sensors are (EM)CCD and sCMOS cameras and Photomultipler tubes (PMTs)
Optical sensors use the photoelectric effect to convert incident photons into electrons
Conversion of photoelectrons to digital information occurs via an analog-digital converter (ADC)
Tradeoffs center around dynamic range and sensitivity

## Slide 3
Objective lenses
https://www.microscope.healthcare.nikon.com/images/mastheads/Imaging-Centers/objective-selector_1920d.jpg

## Slide 4
What factors impact objective lens choice? (5 mins)
https://www.microscope.healthcare.nikon.com/images/mastheads/Imaging-Centers/objective-selector_1920d.jpg

## Slide 5
Magnification
Immersion media
Geometry (eg, dipping?)
Working distance
Aberration correction
Correction collars
N.A.
Cost
Compatibility
Transmission/collection 𝝀
What factors impact objective lens choice?

## Slide 6
Objective lenses
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg, https://www.pinterest.com/pin/227642956134450167/

## Slide 7
Deciphering the objective code
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg

## Slide 8
Objective lenses: magnification
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Determined by focal lengths of internal lenses
1x -> 150x or more
Typical: 4x, 10x, 20x, 40x, 60/63x, 100x
Magnification ≠ Resolution!

## Slide 9
Objective lenses: corrections
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Achieved by internal elements
Plan Apo = planar correction, chromatic correction (achromatic), spherical correction
Cheaper objectives correct for fewer than 3 and/or correct less well

## Slide 10
Objective lenses: immersion
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Immersion choice influences light collection ability (N.A.)
Typical: air, water, glycerol, silicone, oil, oil++
Costs/benefits to consider
  Dry objectives easiest to work with, lowest N.A.
  Water/Oil hardest to work with, highest N.A.

## Slide 11
Objective lenses: N.A.
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
N.A. = numerical aperture
N.A. describes light collection ability of objective
N.A. effectively determines resolution (higher N.A. = better resolution)
Abbe formula
d = 𝝀 / 2 * N.A.
For 520 nm: d = 520 / (2 * 1.45) = ~180 nm

## Slide 12
Objective lenses: working distance
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Typically a function of immersion/N/A.
~100 µm to mm+
Important to consider for volumetric imaging

## Slide 13
Objective lenses: optical path length
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Was finite in older systems, depended on manufacturer
Modern microscopes are “infinity corrected”
Collected light is collimated, converges at infinity
Allows placement of internal optics without introducing aberration

## Slide 14
Objective lenses: coverslip thickness
https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/Objectives/177867/cfi-plan-apo-lambda-d-100x-oil_12689bef50ba806ed3bb0881891e8655.jpg
Typically 0.17 mm (#1.5)
#1 = thinner (0.15 mm)
#2 = thicker (0.22 mm)
Coverslips have a “tolerance” of thicknesses (eg 0.16–0.19 mm for #1.5)
“High tolerance” glass = less deviation
Purpose of correction collar

## Slide 15: µCourse video on PSFs
https://www.youtube.com/watch?v=Tkc_GOCjx7E

## Slide 16
How do we QC an objective?
https://www.google.com/url?sa=i&url=https%3A%2F%2Fwww.researchgate.net%2Ffigure%2FPoint-spread-function-of-a-wide-field-microscope-Different-representations-can-be-used_fig4_288823823&psig=AOvVaw0vynZnpQVxZadEFd3C2EHt&ust=1651597254907000&source=images&cd=vfe&ved=0CAwQjRxqFwoTCJC-vrGlwfcCFQAAAAAdAAAAABAD, https://www.thermofisher.com/order/catalog/product/T7279?ef_id=CjwKCAjwgr6TBhAGEiwA3aVuIUnKf56Ublkjpd5MONdtz-dR4jxdMs97Nz2LJB-4qaZZJjN8RApQihoCKU0QAvD_BwE:G:s&s_kwcid=AL!3652!3!205117794361!!!g!!&cid=bid_pca_ift_r01_co_cp1359_pjt0000_bid00000_0se_gaw_dy_pur_con&gclid=CjwKCAjwgr6TBhAGEiwA3aVuIUnKf56Ublkjpd5MONdtz-dR4jxdMs97Nz2LJB-4qaZZJjN8RApQihoCKU0QAvD_BwE

## Slide 17
Point Spread Functions (PSFs)
https://upload.wikimedia.org/wikipedia/commons/thumb/c/c2/Convolution_Illustrated_eng.png/270px-Convolution_Illustrated_eng.png, https://assets.thermofisher.com/TFS-Assets/BID/product-images/bender_bag_with_one_blue_vial.jpg-650.jpg
Microscopy images are a convolution of real light sources with the PSF of the optical system
PSFs are governed by diffraction
PSFs let us compare performance to theory, across t

## Slide 18
PSFs and 3D imaging
https://www.google.com/url?sa=i&url=https%3A%2F%2Fwww.researchgate.net%2Ffigure%2FPoint-spread-function-of-a-wide-field-microscope-Different-representations-can-be-used_fig4_288823823&psig=AOvVaw0vynZnpQVxZadEFd3C2EHt&ust=1651597254907000&source=images&cd=vfe&ved=0CAwQjRxqFwoTCJC-vrGlwfcCFQAAAAAdAAAAABAD
Microscopy systems are anisotropic (XY ≠ Z)
Lateral (XY): d = 𝝀 / 2 * N.A
Axial (Z): d = 2𝝀 / N.A2
For N.A. 1.45, 520 nm:
  dXY = ~180 nm
  dZ = ~495 nm
Resolution also impacted by aberration (eg, depth)

## Slide 19
Confocal 101
https://static3.olympus-lifescience.com/data/olympusmicro/primer/techniques/confocal/images/confocalintrofigure2.jpg?rev=39B2
Pinhole used to block out of focus light
Pinhole size matched to “Airy disc” (XY PSF)
Confocality = having the same focus

## Slide 20
Selecting confocal light
https://www.photometrics.com/wp-content/uploads/2019/10/Spinning-Disk-Confocal-Microscopy-AppNote.pdf, pg 2

## Slide 21
Spinning disc confocal
https://www.google.com/url?sa=i&url=https%3A%2F%2Fzeiss-campus.magnet.fsu.edu%2Farticles%2Fspinningdisk%2Fintroduction.html&psig=AOvVaw1HYx6BR1qmLhVcactLqzLd&ust=1651598362705000&source=images&cd=vfe&ved=0CAwQjRxqFwoTCIDTwMOpwfcCFQAAAAAdAAAAABAD
An array of pinholes on a disc is rotated at high speed
A second microlens disc can increase excitation efficiency
Allows the use of an EMCCD/sCMOS
Much faster, more sensitive, gentler
Pinhole size fixed

## Slide 22
Challenge 1: Which objective?
http://www.iscopecalc.com/
Go to iscopecalc.com, start with these options:
Which objective do we use of the following?
1: 60x N.A. 1.45 oil 2: 60x N.A. 1.27 water 3: 100x N.A. 1.49 oil

## Slide 23
Challenge 2: Which objective?
http://www.iscopecalc.com/
Go to iscopecalc.com, start with these options:
Which objective do we use of the following?
1: 60x N.A. 1.45 oil 2: 40x N.A. 1.2 water 3: 100x N.A. 1.49 oil

## Slide 24
Challenge 3: Which objective?
http://www.iscopecalc.com/
Go to iscopecalc.com, start with these options:
Identify the ideal objective using the iscopecalc sliders
Go to https://www.microscope.healthcare.nikon.com/products/optics/selector and pick out the best fit you can find