## Slide 1: Fluorescent Labels
Genome 575: Fundamentals of Biomedical Instrumentation
Brian Beliveau

## Slide 2
Recap from last time
Optical systems typically have a defined number of channels
The number of channels is bounded by several factors including label availability and transmission/detection efficiencies
Filters and dichroic mirrors are typically short pass, band pass, or long pass
Adding more channels -> reduced information you can get from any single channel

## Slide 3
On filters…
https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg
It’s complicated!
Let’s walk through some real use cases

## Slide 4
The Beliveau Lab microscopes
https://www.nikon.com/news/2016/img/pic_161109_01.jpg, https://accu-scope.com/media/catalog/product/cache/1/thumbnail/1200x/17f82f742ffe127f42dca9de82fb58b1/6/0/600-3232.jpg, https://cdn.shopify.com/s/files/1/1594/4815/products/061_1200x.jpg?v=1618498984, https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/sites/default/files/individual-filters-small.png
6 filter cubes
7 individual (Em) filters

## Slide 5
Use case 1: widefield LED imaging
https://cms.lumencor.com/system/uploads/fae/image/asset/565/xl_Lumencor_SPECTRAX_ProductGallery-5_2021.jpg, https://cms.lumencor.com/system/uploads/fae/file/asset/30/SPECTRA_X_Light_Engine.pdf?_gl=1*h2z2ow*_ga*OTc4OTE4MzE0LjE2NDk5NjUxNDk.*_ga_364GJZVK77*MTY1MDY2MjAzOS4xLjAuMTY1MDY2MjA1MS40OA..
7 -line LED system
395/25
438/29
474/28
510/25
555/28
640/30
730/40
Goal: Speed

## Slide 6
Use case 1: widefield LED imaging
https://cms.lumencor.com/system/uploads/fae/image/asset/565/xl_Lumencor_SPECTRAX_ProductGallery-5_2021.jpg, https://www.chroma.com/products/sets/89404-et-391-32-479-33-554-24-638-31-multi-led-set-with-single-band-emitters, https://cdn.shopify.com/s/files/1/1594/4815/products/061_1200x.jpg?v=1618498984, https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/sites/default/files/individual-filters-small
395/25
438/29
474/28
510/25
555/28
640/30
730/40
7 -line LED system
Goal: Speed

## Slide 7
Use case 1: widefield LED imaging
https://cms.lumencor.com/system/uploads/fae/image/asset/565/xl_Lumencor_SPECTRAX_ProductGallery-5_2021.jpg, https://www.chroma.com/products/sets/89404-et-391-32-479-33-554-24-638-31-multi-led-set-with-single-band-emitters, https://cdn.shopify.com/s/files/1/1594/4815/products/061_1200x.jpg?v=1618498984, https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/sites/default/files/individual-filters-small
395/25
438/29
510/25
555/28
640/30
730/40
7 -line LED system
Goal: Speed

## Slide 8
Use case 2: Super Resolution
https://i.ytimg.com/vi/y4dJOFUVk30/maxresdefault.jpg, https://d33b8x22mym97j.cloudfront.net/production/imager/productphotos/827/filter-cubes-main_ec7d1033075031a0254a7350f4e87eb4.jpg, https://www.chroma.com/sites/default/files/individual-filters-small, https://www.chroma.com/spectra-viewer?parts=25787%2C25416%2C25506#tabs-spectra_viewer_plot-left-3
532 nm
Goal: Photons

## Slide 9
What makes a good label? (5 mins)
https://www.idtdna.com/site/Catalog/Modifications/Dyes, http://www.chemspider.com/Chemical-Structure.23107060.html

## Slide 10
What makes a good label?
https://www.idtdna.com/site/Catalog/Modifications/Dyes, http://www.chemspider.com/Chemical-Structure.23107060.html
Brightness
Photostability
Duty cycle
Bio-orthogonality / inertness
Chemical stability (pH)
Size
Monovalency
Cost / availability
Thermal stability
Stokes shift
Em/Ex wavelengths
Solubility

## Slide 11
How do we evaluate labels?
https://www.atto-tec.com/ATTO-565.html?language=en

## Slide 12
Measuring brightness: molar absorbance coeffcient (ε)
https://resources.perkinelmer.com/lab-solutions/resources/docs/APP_Determination_of_Relative_FluorescenceQuantum_Yields_using_FL6500_Fluorescence_Spect.pdf, https://www.edinst.com/wp-content/uploads/2019/08/BLL-1.png, https://www.edinst.com/wp-content/uploads/2019/08/BLL-E4.png

## Slide 13
Measuring brightness: Quantum Yield (ɸ)
https://www.edinst.com/wp-content/uploads/2019/08/BLL-1.png , https://www.researchgate.net/profile/Henri-Pierre-Lassalle/publication/30513452/figure/fig1/AS:913119526273024@1594716068206/Jablonski-diagram-where-IC-stands-for-internal-conversion-ICS-for-intersystem.png
S0
S1
Sn
ɸ = emitted photons / absorbed photons

## Slide 14
Measuring brightness
https://www.atto-tec.com/ATTO-565.html?language=en
Brightness = ε * ɸ

## Slide 15
Labels: organic dyes
https://www.atto-tec.com/ATTO-565.html?language=en, https://www.idtdna.com/site/Catalog/Modifications/Dyes
Hydrocarbon-based
~1 nm in size
Typically the best bang for buck in terms of brightness/stability, sample compatibility

## Slide 16
Labels: Fluorescent proteins
http://www.chemspider.com/Chemical-Structure.23107060.html, https://zeiss-campus.magnet.fsu.edu/articles/probes/images/anthozoafpintrofigure1.jpg
Genetically encoded
~5–10 nm in size
Typically 2–10x dimmer than organic dyes
Often multimeric (aggregation artifacts)

## Slide 17
Labels: Quantum Dots
https://www.displayninja.com/wp-content/uploads/2019/06/what-is-qled-technology.jpg
Have complex structure
~20–50 nm in size after coatings
Typically 2–10x brighter than organic dyes
Difficult to use in cellular systems*

## Slide 18
Class activity 1: Dye selection
Go to https://www.fpbase.org/spectra/ and add the filters and light sources below
Select one ATTO or Alexa Fluor dye for the “Cyan”, “GreenYellow”, and “Red” channels. Add “DAPI” for “Violet”
Aim to balance excitation and emission efficiencies

## Slide 19
Class activity 2: FP slection
Go to https://www.fpbase.org/spectra/ and add the filters and light sources below
Select one fluorescent protein for the “Cyan”, “GreenYellow”, and “Red” channels. Add “DAPI” for “Violet”.
Aim to balance excitation and emission efficiencies